Nursing Education and Development

EkoSonic® Endovascular System Skills Checklist

| CRITICAL DECISION POINTS |  | MET | NOT MET |
| --- | --- | --- | --- |
| 1. | The UPS/battery should be plugged in to a working outlet at all times except during transport. |  |  |
| 2. | Catheter cables (gray and black) are connected to the same Control Unit’s (CU) Connector Interface Cable (CIC). |  |  |
| 3. | Ultrasound power is auto regulated by the EkoSonic® CU. |  |  |
| 4. | Drug (Lytic) and coolant (Normal/Heparinized Saline) are being infused through hospital infusion pumps. The EkoSonic® System in no way controls drug infusion. |  |  |
| 5. | The CU alarms can be silenced by pushing the upper right hand soft key beside the Alarm Silence Icon. If ultrasound is stopped, it must be re-started by pressing the green START button. If this is unsuccessful then proceed to Item 9 below. |  |  |
| 6. | Never aspirate from the drug lumen. This will clog the micro-pores with blood and render the catheter useless. |  |  |
| 7. | Proper functioning of the EkoSonic® System is confirmed by a flashing yellow light below the display & immediately to the left of the green START button. |  |  |
| 8. | If thrombolytic drug is discontinued: - Keep coolant running at a maintenance level to avoid the catheter clotting off. - Obtain physician’s order to infuse normal saline through the drug port at a maintenance level in case thrombolytic drug is restarted. - Ultrasound should not remain on unless fluid is infusing through the drug (min 5 ml/hr) and coolant ports (min 35 ml/hr). - Inform the physician ultrasound should be turned off without the minimum infusion rates. Press the orange STOP button. Locate the rocker switch on the back of the Control Unit. Turn off the Control Unit. |  |  |

| CRITICAL ELEMENTS |  | MET | NOT MET |
| --- | --- | --- | --- |
| 1. | Plug the UPS/battery into an AC outlet and ensure that the CU is plugged into the UPS/battery and the UPS/battery is turned ON (left hand button). |  |  |
| 2. | Confirm that the CIC is connected to the front of the EkoSonic® Control Unit. |  |  |
| 3. | Connect the catheter cables (black and gray) to the CIC according to the color code and symbol. Note: the red circles around the MSD and IDDC connections on the CU display will disappear if the connection is good. |  |  |
| 4. | Verify coolant and thrombolytic infusion rates for accuracy. Verify that the Normal/Heparinized Saline is connected to the coolant port of the EKOS® IDDC and the thrombolytic is connected to the drug port of the EKOS® IDDC. |  |  |
| 5. | Infusion to the side arm of the sheath is at the discretion of the physician. Check MD orders. |  |  |
| 6. | Push the green START button on the CU and ensure the yellow light is blinking in the left lower corner of the unit. The CU display may be left showing the default Home screen or may be switched to the Power Graph screen as desired. |  |  |

| CRITICAL ELEMENTS | MET | NOT MET |
| --- | --- | --- |

| 7. | Push the yellow STOP button when performing Doppler checks or Echocardiograms. An alarm will occur after 5 minutes as a reminder to restart the ultrasound. Pushing the Alarm Silence button will reset the 5 minute timer. Resume ultrasound after the Doppler or Echo is finished by pressing the green START button. Ensure the yellow light is blinking. |  |  |
| --- | --- | --- | --- |
| 8. | Silence the battery alarm by depressing the MUTE button (right most one) located on the UPS/battery after unplugging the unit to transport the patient. |  |  |
| 9. | Troubleshooting: - If a CU alarm sounds it can be muted by pushing the upper right hand soft key beside the Alarm Silence Icon. - Confirm that the EkoSonic® Display does not show any red circles indicating a disconnected cable. Reconnect any cable showing a circle and confirm that it goes away. Push the green START button. - If the alarm condition persists then identify the troubleshooting icon that is flashing in the bottom right hand portion of the CU display. - Contact EKOS support by calling the number located on top of the EkoSonic® control unit. (1-888-356-7435) - If unable to resume ultrasound therapy after contacting EKOS then call the Interventional Physician. |  |  |
| 10. | Removal of the device at bedside: a. Turn off the ultrasound, disengage black connector clip, disconnect the grey and black connectors from the connector interface cable b. Turn off the stopcocks from the drug and coolant ports to the patient. c. Don PPI and place the patient flat d. Have the patient exhale and hold breathe if able e. Stabilize the introducer sheath with one hand and then pull the EKOS® catheter out in a steady motion. If resistance is noted, stop and notify physician. f. If sheaths are to be discontinued follow the hospital policy. |  |  |

|  | ## Passed |  | Needs to repeat |
| --- | --- | --- | --- |

### Validated by: ________________________________ Date:  ____________

Forms: Ekos –  Nursing

This checklist is intended as an example to demonstrate the type of checklist you may
wish to implement in your clinical practice. Any checklist you implement should reflect your actual clinical practice at your facility. EKOS makes no recommendations or representations about the content of this sample checklist and the ultimate responsibility for such a checklist rests with the clinical practice.